1. INTRODUCTION

The 4 in 1 digital multi- Multi-Function Environment Meter has been designed to combine the functions of Sound Level Meter, Light Meter, Humidity Meter, Temperature Meter.

It is an ideal Multi-Function Environment Meter Instrument with scores of practical applications for professional and home use.

The Sound Level function can be used to measure noise in factories, schools, offices, airports, home, etc., checking acoustics of studios, auditoriums and hi-fi installations.

The Light function is used to measure illuminance in the field. It is fully cosine corrected for the angular incidence of light. The light sensitive component used in the meter is a very Stable, long life silicon diode.

The Humidity/Temperature is for use a humidity/

semiconductor sensor and K type thermocouple. This operations manual contains general information and specification

- FEATURES

- Optically isolated PC interface-optional RS-232 cable and windows® 95/98/2000 compatible software allows user to collect display, plot and save data

- 4 functions measure Sound level, Light, Humidity, Temperature.

- 4000 count large LCD display with units of Lux, oC,  %RH and dB indication.

- Easy to use with single function switch operating, pocket size and light weight.

- Sound level measures from Lo: 35dB to 100dB/ Hi: 65dB to 130dB for A weighting checking with 0.1dB resolution.

- Light measuring levers ranging from 0.01 lux to 40,000 lux.

- Humidity measurement from 25%RH to 95%RH with 0.1%RH resolution and fast time response.

### 3. SPECIFICATIONS

Display: 4000 counts LCD display with function of Lux, oC, oF, %RH and dB indication.

Polarity: Automatic, (-) negative polarity indication.

Over-range: “OL” mark indication.

Low battery indication: The “BAT” is displayed when the battery voltage drops below the operating level.

Measurement rate: 1.5 times per second, nominal.

Storage temperature: -10 oC to 60 oC (14 oF to 140 oF) at < 80 % relative humidity

Power: One standard 9V, NEDA1604 or 6F22 battery.

Dimensions: 121.5 (H) x 60.6 (W) x 40 (D) mm

Weight: Approx.: 280g including holster.

Sound Level

Measurement range:

Lo: 35dB -100dB;  Hi: 65dB~130dB

Resolution: 0.1dB

Typical instrument frequency range: 30Hz-10KHz

Frequency Weighting: A –weighting

Time Weighting: Fast

Accuracy: +3.5 dB at 94 dB sound level,

1KHZ sine wave.

Microphone: Electric condenser microphone.

Light

Measuring Range:40, 400, 4000, 40,000lux

(40,000lux range reading x10)

Overrate Display: Highest digit of “OL” is displayed.

Accuracy: +5% rdg + 10 dgts (calibrated to standard incandescent lamp at color temperature 2856 k ) .

Repeatability: + 2%.

Temperature Characteristic: +0.1% / oC.

Photo detector: One silicon photo diode with filter.

Humidity/Temperature(Auto-ranging)

Measurement Range:

Humidity 25%~95%RH

Temperature: -20.0℃-+400.0℃, -20℃-+750℃,

-4.0℉-+400.0℉, -4℉-+1400℉

Resolution : 0.1% RH, 0.1℃or 1℃/ 0.1℉or1℉.

Accuracy (after calibration):

Humidity: + 5%RH (at 25℃ , 35%～95% RH ).

Response time of the humidity sensor: approx. 6 min.

Temperature:

+3%rdg + 2℃ (at -20℃~+400.0℃)

+3.5%rdg + 2℃ (at -20℃~+750℃)

+3%rdg + 2℃ (at -4.0℉~+400.0℉)

+3.5%rdg + 2℃ (at -4℉~+1400℉)

Input Protection: 60V dc or 24V ac rms.

4. PANEL DESCRIPTION

1.  LCD display: 4000 count LCD display with units of Lux, x10 Lux, oC, oF, %RH, dB and low battery “BAT” indication.

- Power / Function / Range Switch: Turn power on (or    off) and select measurement function and ranges.

- Microphone: Electric condenser microphone inside.

- Photo Detector: Long life silicon photo diode inside.

- Humidity & Temperature: Humidity Sensor and  Semiconductor Sensor inside.

- Temperature Terminal: Insert the temperature probe in this terminal.

- RS-232 Terminal : Computer Interface Transport Terminal

** 8. RS-232 wires: Connect the RS-232 wires to the meter’s “RS-232” interface

5. OPERATING INSTRUCTION

Measuring Sound Level

- Turn the Power/function/range Switch to “dB” position.

- Remove the meter and face the microphone to sound source in a horizontal position.

- The A -weighting curve is nearly uniform over the frequency range from 30 to 10,000Hz, thus giving an indication of overall Sound level.

- The Fast response is suitable to measure shout bursts and peak values from sound source.

- The sound level will be displayed.

- Note: Strong wind (over 10m/sec.) striking the microphone can cause misreading for measurement in windy locations, a windscreen should be used in front of microphone.

Measuring Light

- Turn the Power/function/range Switch to select the “lux” scale and set the range to desired (“lux” or “x10 lux”) range.

- Remove the meter and face the photo detector to light source in a horizontal position.

- Read the illuminance nominal from the LCD display.

- Over-range: If the instrument only display one “OL” in the M.S.D. the input signal is too strong, and a higher range should be selected.

- When the measurement is completed. Replace the photo detector from the light source.

- Spectral sensitivity characteristic: To the detector, the applied photo diode with filters makes the spectral sensitivity characteristic almost meet C.I.E. (International Commission on Illumination) photopia curve V () as the following chart described.

100% (Relative Sensitivity)    Spectral Sensitivity

|  |  |  |  |  |  |  | V( ) |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  |

					

Wavelengthnm)

- Recommended Illumination:

Locations						     Lux

*Office

Conference, Reception room.	 200	~  750

Clerical work					 700	~ 1,500

Typing drafting				 1000 ~ 2,000

*Factory

Packing work, Entrance passage	 150	~  300

Visual work at production line	 300	~  750

Inspection work				 750	~ 1,500

Electronic parts assembly line	 1500	~ 3,000

*Hotel

Public room, Cloakroom		 100	~  200

Reception, Cashier			 200	~ 1,000

*Store

Indoors Stairs Corridor		 	 150	~  200

Show window, Packing table	 750	~ 1,500

Forefront of show window		 1500	~ 3,000

*Hospital

Sickroom, Warehouse			  100	~  200

Medical Examination room		  300	~  750

Operating room

Emergency Treatment			  750	~ 1,500

*School

Auditorium, Indoor Gymnasium	  100	~  300

Class room				 	  200	~  750

Laboratory Library Drafting room  500	~ 1,500

Measuring Humidity/Temperature

1. Humidity Measurement:

① Set the Power/function/range Switch to “%RH” position .

②  Then the display will show the humidity reading value (%RH) directly.

③ When the tested environment humidity value changed. It need to a few minutes to get the stable “%RH” reading.

Warning:

Don’t expose the humidity sensor to direct sunlight.

Don’t touch or manipulate the humidity sensor.

2. Temperature Measurement:

① Set the Power/function/range Switch to “0.1 oC or

1 oC and 0.1℉ or 1 ℉” position.

② Then the display will show the environment

temperature reading value (oC/℉) directly.

③Insert the temperature probe into the K-type thermocouple socket.

④ Touch the end of the temperature sensor to the area or surface of the object to be measured. The display will show the temperature reading value (oC/℉) directly.

Warning:

When function switch on temperature “0.1 oC or 1 oC and 0.1℉ or 1 ℉”range, Never attempt a voltage measurement with the test leads inserted into the K-type thermocouple socket.

You might be injured or damage the meter.

#### RS-232 INTERFACE MEASUREMENTS

- Connect the RS-232 wires to the meter’s “RS-232” interface and PC “COM” port., the indicator “RS232” will be appear in the display.

- PC will display the meter’s measuring model.

Note: First of all please read the “README” file in the attached soft disk and then comply with it set the SET UP Program from soft disk up to your PC.

6. MAINTENANCE

#### Battery Replacement

If the sign “BAT” appears on the LCD display, it indicates that the battery should be replaced. Remove screws on the back cover and open the case. Replace the exhausted battery with new batteries. (1 x 9V battery NEDA 1604, 6F22 or equivalent)

# INSTRUCTION MANUAL

### MODEL 2232

### Multi-Function Environment Meter

With PC Interface

###### 4 IN 1

SOUND LEVEL

LIGHT

HUMIDITY

TEMPERATURE

##### TABLET OF CONTENTS

TITLE

- INTRODUCTION

- FEATURES

- SPECIFICATIONS

- PANEL DESCRIPTION

- OPERATING INSTRUCTION

- MAINTENANCE